## System Evaluation
| Participant 1 | Participant 2 | Participant 3 | Participant 4 | Participant 5 | Mean | STD
The system is easy to use. | 2 | 4 | 4 | 5 | 4 | 3.8 | 1.1
The system provides all the motions required to perform ablation. | 3 | 4 | 4 | 5 | 4 | 4 | 0.7
The system was reliable and consistent during the procedure. | 3 | 4 | 4 | 5 | 5 | 4.2 | 0.8
The system easily integrates with existing equipment in the X-MRI lab. | 4 | 3 | 4 | 5 | 5 | 4.2 | 0.8
The visual feedback, such as the LEDs, simulation, and video feedback, provide sufficient information on what is happening inside the lab. | 3 | 4 | 4 | 4 | 5 | 4 | 0.7
The simulation of the robot alone, provides sufficient information on what is happening inside the lab. | 2 | 3 | 4 | 4 | 4 | 3.4 | 0.9
The video feedback from the camera, provides sufficient information on what is happening inside the lab. | 3 | 3 | 4 | 3 | 5 | 3.6 | 0.9
There were no limitations or challenges encountered during the procedure. | 1 | 4 | 4 | 3 | 4 | 3.2 | 1.3

## Maneuvers Evaluation
| Speed | [6 empty] | Sensitivity
| Participant 1 | Participant 2 | Participant 3 | Participant 4 | Participant 5 | Mean | STD | Participant 1 | Participant 2 | Participant 3 | Participant 4 | Participant 5 | Mean | STD
Insertion. | 4 | 4 | 4 | 5 | 4 | 4.2 | 0.4 | 4 | 4 | 4 | 5 | 5 | 4.4 | 0.5
Advancement. | 5 | 4 | 4 | 4 | 4 | 4.2 | 0.4 | 5 | 4 | 4 | 5 | 5 | 4.6 | 0.5
Handle rotation. | 2 | 3 | 4 | 4 | 3 | 3.2 | 0.8 | 2 | 3 | 3 | 5 | 3 | 3.2 | 1.1
Catheter tip flexion and extension. | 4 | 4 | 4 | 4 | 3 | 3.8 | 0.4 | 1 | 4 | 3 | 4 | 3 | 3 | 1.2
Steering. | 3 | 3 | 4 | 4 | 3 | 3.4 | 0.5 | 3 | 3 | 4 | 4 | 3 | 3.4 | 0.5
Withdrawal. | 5 | 4 | 4 | 5 | 4 | 4.4 | 0.5 | 5 | 4 | 4 | 5 | 5 | 4.6 | 0.5

## Joystick Evaluation
| Participant 1 | Participant 2 | Participant 3 | Participant 4 | Participant 5 | Mean | STD
The interaction with the robot is natural, i.e., left hand for linear translation, both hands for rotation, and right hand for tip deflection. | 4 | 4 | 4 | 4 | 5 | 4.2 | 0.4
The joystick controller provides intuitive control. | 3 | 3 | 3 | 4 | 4 | 3.4 | 0.5
The joystick interface is easy to navigate. | 3 | 4 | 4 | 4 | 4 | 3.8 | 0.4
The joystick controller is comfortable to hold during the procedure. | 4 | 4 | 4 | 5 | 5 | 4.4 | 0.5
The joystick is responsive. | 3 | 4 | 4 | 4 | 3 | 3.6 | 0.5